## Rack requirements
Power supply for sensor: voltage=300 V; current= 4 mA. / Analogue power supply for readout: voltage= 1.2 V; current= 199 mA. / Digital power supply for readout: voltage= 1.2 V; current=1.22 A. / Core power supply for readout: voltage= 1.5 V; current= 943 mA. / Transceiver power supply for readout: voltage= 2.5 V; current= 425 mA.
Sensor
voltage	current			suggestion	required channel	channel	module	crate	unit
[V]	[mA]			[#]	[#]	[#]	[#]	[#]	[U]
500	15			a1520p	176	194	17	3	24
Readout (option A)
analogue		digital
voltage	current	voltage	current	suggestion	required channel	channel	module	crate	unit
[V]	[A]	[V]	[A]	[#]	[#]	[#]	[#]	[#]	[U]
7.0	4	7.0	4	a1517b	636	700	117	8	64
Readout (option B)
analogue + digital
voltage	current			suggestion	required channel	channel	module	crate	unit
[V]	[A]			[#]	[#]	[#]	[#]	[#]	[U]
4.5	6			a1518b	318	350	59	4	32
Transmission
core		transceiver
voltage	current	voltage	current	suggestion	required channel	channel	module	crate	unit
[V]	[A]	[V]	[A]	[#]	[#]	[#]	[#]	[#]	[U]
7.0	4	7.0	4	a1517b	244	269	45	3	24
Total (option A)
					crate	unit
					[#]	[U]
					14	112

## Reference
Part of the system	Comments	Partial		Total
cooling plant (Pixel +Strips)	pump+vacuum pump+plc+controllers	10 kW		40 kW
	Chiller (400VAC tri +N) if not cool water from FAIR	4 kW
	Heaters ((0.3-0.4 kW/ch) + transducers	26 kW
cooling system (x services)	chiller, pump,plc, controllers, transducers	10 kW		10 kW
Part of the system	Comments	Single	Quantity	Total
Power supply Low Voltage	176 digital channels + 176 digital channels	3.4 kW	4	13.6 kW
Pixel readout chips	power extrapolated from Topix3-> final Topix
	DC-DC efficiency ~60%
	cable losses- factor 2 + safety factor 2
	13 kW as raw power ( over exstimated)
	power supply - CAEN SY1527-16 boards
	A1513B board (6ch/board)
Power supply High Voltage	176 sensor channels	3.4 kW	1	3.4 kW
Pixel sensor	CAEN SY1527- 16 boards
	A1510 board (12 ch/board) ( 100 V-limit)
El/opt converter	122 GBT channels	3.4 kW	1	3.4 kW
Pixel part	DC-DC efficiency ~ 60% (50 boards)
	cable losses -factor 2 + safety factor 2
	2.8 kW as raw power
	power supply - CAEN SY1527-16 boards
	A1513B board (6ch/board)
electronics crate	on-detectors electronics, interlocks	3 kW	1	3 kW
pixel part	power supply controllers
electronics crates-counting room	Optical receiver + data concentrator	3 kW	6	18 kW
pixel part
FE Module power supply	254 modules, 15 W/module incl. Mod. Data Concentrator	3.4 kW	6	20.4 kW
strip part	and GBT Interface, 0,6 DCDC-Efficiency
	power supply - CAEN SY1527-16 boards
	A1513B board (6ch/board)

HV power supply	254 sensors, CAEN power supply (see above)	3.4 kW	2	6.8 kW
strip part
El/opt converter	254 GBT channels	3.4 kW	2	6.8 kW
strip part
electronics crate	on-detectors electronics, interlocks	3 kW	1	3 kW
strip part	power supply controllers
electronics crates-counting room	Optical receiver + data concentrator	3 kW	6	18 kW
strip part